## Sheet1
DEVICE | MONOPOLAR (BOVIE) SAFE | MRI SAFE
1. Non-programmable shunt valves | SAFE | SAFE
2. Programmable shunt valves | SAFE | Safe, but check settings after scan
3. Baclofen Pump | YES, minimum of 15cm separation between pump or catheter | Safe, but check functionality after scan
4. Vagus Nerve Stimulator (VNS) | Patient with system already in place: safe if output current parameters are set to zero. Flow through the generator and leads must be minimized. Patient getting new implant: Can use cautery to create pocket. Stop use once generator is on the sterile field. | Safe if area of implant is excluded from radiofrequency field (C7-T8) *Manufacturer guidelines must be strictly followed.
5. Deep Brain Stimulator (DBS) | NOT SAFE | Safe only if device is turned off *Manufacturer guidelines must be strictly followed
6. Intellis Spinal Cord Stimulator | NOT SAFE | Device is to be in 'MRI SAFE' mode. Follow manufacturer guidelines.
7a. Responsive Neural Stimulation (RNS) RNS-320M model | NOT SAFE | Device is to have 'MRI MODE ON' for up to 1.5T MRI. Follow manufacturer guidelines.
7b. Responsive Neural Stimulation (RNS) RNS-300M model | NOT SAFE | NOT SAFE